# Design of a Flow Velocity Meter for Tides in the River

# Using Arduino

Sanya Samaimak1 and Shanin Harnnarong1

**ABSTRACT**

The Royal Irrigation Department is responsible for managing water availability in Thailand. Many instruments are used to read the data to obtain the needed hydrological data. The instruments that it has applied to measure the velocity of the tide are the flow velocity meters A-OTT C31 that compatible with the audio and numerical displayer Z 41-00. They have been used for 30 years, (1988 - 2018).

Design of a flow velocity meter for tides in the river using Arduino was presented in this paper. It was designed to use as a substitute for the audio and numerical displayer Z 41-00 that were broken. The result of the design and the experimentation show the ability of working together with the flow velocity meters A-OTT C31 and the accuracy of the data acquired from this designed instrument are satisfactory.

**Keyword:** The Royal Irrigation Department, Flow Velocity Meter, Arduinos

**Introduction**

The water current velocities are measured by the tools that measure the speed of the tide and they are classified into two types.

**1. Mechanical current meters**

Mechanical current meters are mechanical devices that are the main components. The tools will move when the current flows and there are three types.

1.1 Vertical axis current meters

1.2 Horizontal axis current meters

1.3 Pendulum current meters

**2. Electronic current meters**

Electronic current meters are electronic devices that work primarily on electronic devices. These tools work better than mechanical current meters and there are three types.

2.1 Electromagnetic velocity meters

2.2 Doppler velocity meters

2.3 Optical strobe velocity meters

The tools that have been used by the Irrigation Department are A-OTT C31, the current flow meters, with Z 41-00, audio and numerical indicators. The A-OTT C31 current flow meter is mechanical current meter that is the type of horizontal axis current meter.

1 Department of Instrumentation Engineering, Faculty of Engineering, Rajamangala University of Technology Rattanakosin, 96 Moo 3 Salaya Phutthamonthon NakhonPathom 73170, Thailand

\*(Corresponding author, e-mail) : sanya.sam@rmutr.ac.th

**Materials and methods**

**1. Signal form A-OTT C31**

|  |  |
| --- | --- |
|  |  |
| **Figure 1** A-OTT C31 | **Figure 2** A-OTT C31 Signal |

**2. Block diagram**



**Figure 3** Block Diagram

**3. Flowchart**



**Figure 4** Flowchart

**Results and discussions**

**1. Testing at 20 Hz, 200 Hz, 1 kHz and 2 kHz**

|  |  |
| --- | --- |
|  |  |
| **Figure 5** Test bench. |

**Table 1** Test results at 20 Hz

|  |  |  |
| --- | --- | --- |
| **Oscilloscope** | **Serial Monitor** | **Summary** |
|  |  |  |

**Table 2** Test results at 200 Hz

|  |  |  |
| --- | --- | --- |
| **Oscilloscope** | **Serial Monitor** | **Summary** |
|  |  |  |

**Table 3** Test results at 1 kHz

|  |  |  |
| --- | --- | --- |
| **Oscilloscope** | **Serial Monitor** | **Summary** |
|  |  |  |

**Table 4** Test results at 2 kHz

|  |  |  |
| --- | --- | --- |
| **Oscilloscope** | **Serial Monitor** | **Summary** |
|  |  |  |

**Conclusions**

The experimental results show that this designed prototype can read four tested areas of the frequency range properly well. The application of this prototype will be applied to the A-OTT c31, the current flow meter, for replacing the damaged Z 41-00, the counter which is used to display as an audio and numerical output. It can be done simply by connecting only two wires. The prototype of this research will be developed as a tool to be used in fieldwork in the future.

**Acknowledgements**

This research was supported by the Institute of Research and Development, Rajamangala University of Technology Rattanakosin. In addition, the authors would like to thank Mr.Voravot Boontong, Hydrologic Instrument Standards Branch Chief, working in Hydrology Division, Bureau of Water Management and Hydrology for all his help and suggestions. Thankfulness to the supporters : Mr.Jetsdaporn Satansup, Mr.Ruangsimon Jamkrajang and Mr.Pawaritsorn Chaisong for all their support throughout the period of this research.

**References**

OTT Hydromet. 2017. A-OTT C31 [Online]. Available: http://www.ott.com/products.

Arduino Language Reference [Online]. https://www.arduino.cc/reference/en.

Makarn, Eakachai. 2009. Learn Understand and Use the AVR Micro controller family with

Arduino. Bangkok: ETT Co., Ltd. (in Thai)

เอกชัย มะการ. 2552. เรียนรู้ เข้าใจ ใช้งานไมโครคอนโทรลเลอร์ตระกูล AVR ด้วย Arduino. กรุงเทพมหานคร:

บริษัท อีทีที จำกัด.

Polpananavee, Pramoht. 2011. Principle of Water Flow calculation through Irrigation

Structures. Bangkok: Regional Irrigation Office 8. (in Thai)

ปราโมท พลพณะนาวี. 2554. หลักการคำนวณน้ำผ่านอาคารชลประทาน. กรุงเทพมหานคร:

สำนักกรมชลประทานที่ 8.

Leevajanakul, Kirati. 2000. Hydrology. Bangkok: Rangsit University Press. (in Thai)

กีรติ ลีวัจนกุล. 2543. อุทกวิทยา. กรุงเทพมหานคร: สำนักพิมพ์มหาวิทยาลัยรังสิต